# Gas Permeability of Saint‑Gobain Sani‑Tech Ultra Silicone Tubing and Best-Available Measured Analog Data

## Executive summary

Publicly accessible Saint‑Gobain documentation for Sani‑Tech® Ultra silicone tubing (Ultra‑C / Ultra‑65) provides dimensions, part-number decoding, and mechanical properties, but does not publish measured O₂/CO₂/N₂ permeability/permeation values (nor an ASTM D1434 report) in the open datasheet. [1]

The Saint‑Gobain product page explicitly indicates that technical dossiers / validation guides are “gated documentation” (i.e., not publicly downloadable without registered access), which is the most likely location for any internal or third‑party gas transport characterization (if it exists). [2]

Because ULTRA‑C‑125‑1F / ULTRA‑C‑125‑1H permeability values were not found as product-specific public data, the best authoritative, measured substitutes located are: - A widely cited primary measurement set for dimethyl silicone rubber (with silica filler) at ~25 °C, reporting N₂ ≈ 28×10⁻⁹, O₂ ≈ 60×10⁻⁹, CO₂ ≈ 325×10⁻⁹ in the classic composite permeability units (cc gas·cm)/(s·cm²·cmHg·Δp). [3]
- A second measured dataset (same paper) showing similar magnitudes and explicitly discussing temperature dependence and the role of diffusion/solubility (P = D×S). [4]
- A platinum‑cured PDMS tube vendor page with permeability values reported in a “10⁻⁹ Barrer-style” convention (numerically similar scale to the classic silicone-rubber datasets, but not Saint‑Gobain ULTRA-specific). [5]

Using the wall-code 1 thickness = 0.79 mm (per Saint‑Gobain’s PN decoding) and a planar steady-state approximation, these literature-grade permeability values imply order-of-magnitude fluxes at Δp = 1 atm of roughly N₂ ~0.25 cm³(STP)/m²/s, O₂ ~0.53 cm³(STP)/m²/s, CO₂ ~2.9 cm³(STP)/m²/s (divide by 10 for Δp = 0.1 atm). (Method and assumptions detailed below.) [6]

flowchart TD
  A[Goal: measured O2/CO2/N2 permeability for ULTRA-C-125-1F/1H] --> B[Search Saint-Gobain pages & datasheets]
  B --> C{Public O2/CO2/N2 permeability values found?}
  C -->|No| D[Extract PN decoding + geometry from datasheet]
  D --> E[Search peer-reviewed / primary measured silicone rubber & PDMS data]
  E --> F[Robb 1968: silicone rubber permeability tables]
  E --> G[Other corroborating sources: Arlon paper; PDMS tube vendor]
  F --> H[Convert units & compute flux for wall = 0.79 mm at Δp=1 atm and 0.1 atm]
  G --> H

## Product definition and geometry for Ultra‑C‑125‑1F / Ultra‑C‑125‑1H

### What “Ultra‑C‑125‑1F/1H” encodes

Saint‑Gobain’s datasheet defines a part number structure ULTRA‑X1‑XXX‑XX2, where: - X1 indicates hardness: C = 50 Shore A, 65 = 65 Shore A - XXX is ID in inches (e.g., 125 → 0.125 in) - Wall size code: 1 = 1/32" = 0.79 mm - Coil/roll: F = 50 ft, H = 100 ft, and 25 ft has no designation [7]

Saint‑Gobain’s product page also characterizes Sani‑Tech Ultra as platinum‑cured tubing available in 50 or 65 Shore A options. [2]

### Dimensions for the Ultra‑C‑125 wall-code 1 variant

From the standard sizes table: - ULTRA‑C‑125‑1: ID 0.125 in (3.2 mm), OD 0.188 in (4.8 mm), wall 0.032 in (0.8 mm) [8]
This is consistent with the PN “wall code 1 = 0.79 mm” mapping. [7]

Working thickness assumption for calculations: wall-code 1 → t = 0.79 mm = 0.00079 m. [7]

## Search outcome for product-specific O₂, CO₂, N₂ permeability data

### Saint‑Gobain publicly available documents

The open Saint‑Gobain datasheet provides: - a dimensional table of part numbers, ID/OD/wall, and burst pressure - part number decoding including coil lengths and wall-code mapping - mechanical properties testing context (room temperature, sample geometry) [9]

However, no O₂/CO₂/N₂ permeability coefficients (in Barrer or SI) and no gas transmission test report (e.g., ASTM D1434) are included in the public datasheet. [9]

### Why you may not see it publicly

The Saint‑Gobain product page indicates a “Resource Center” containing Technical Dossiers and Validation Guides as gated documentation. If Saint‑Gobain (or a customer-facing test lab) has performed gas permeability characterization on ULTRA formulations, this gated set is the most plausible location. [2]

## Best-available measured permeability data for comparable silicone/PDMS materials

### Primary measured silicone rubber data (closest “material class” match)

W. L. Robb (1968) reports measured permeability constants for dimethyl silicone rubber with silica filler (explicitly: dimethyl siloxane + 33 wt% silica filler) and gives permeability at 25 °C, expressed as:

Pr × 10⁹ = (cc gas at 1 atm, 25 °C)·cm / (s·cm²·cmHg·Δp) [3]

From Robb’s Table 2 (filled dimethyl silicone rubber): - N₂: 28
- O₂: 60
- CO₂: 325 [3]

Robb details the experimental approach (two-bulb apparatus evacuated to <10⁻³ mmHg, gas admitted at known pressure, pressure buildup measured on the low-pressure side; reported volumes referenced to 1 atm and 25 °C). [10]

Robb also states the standard solution–diffusion framing Pr = D × S (permeability as the product of diffusivity and solubility), and provides explicit D and S values for key gases in silicone rubber. [4]

### Sensitivity to formulation (important for “ULTRA vs literature” gap)

Robb reports that across 0–60 wt% silica filler, the O₂ permeability only varied from ~70×10⁻⁹ down to ~40×10⁻⁹ in the same composite units—i.e., formulation changes can move permeability by ~±30–40% even within “dimethyl silicone rubber” families. [11]
This is directly relevant because tubing compounds often include fillers and processing-induced morphology that differ from thin membrane films.

### Corroborating tabulation with D, S, and gas set (not ULTRA-specific)

A technical paper by Zhang & Cloud (Arlon, reprinted from SAMPE 2006) reproduces a table of dimethylsilicone rubber gas permeability for CO₂, N₂, O₂ alongside diffusivity and solubility values, using the same classic composite permeability unit scale (×10⁹). [12]
It also documents the use of ASTM D1434-82 (Procedure V, volumetric) for permeability testing in their experimental work (though their own reported test gas in that section is helium). [13]

### Platinum-cured PDMS tube vendor values (measured, but not Saint‑Gobain ULTRA)

A vendor page for platinum-cured silicone (PDMS) tube lists “Gas Permeability Rates” for O₂/CO₂/N₂ and specifies the unit convention as:
“In barrers 10⁻⁹ (cc gas(RTP)cm)/(sec cm² cmHg ΔP)” [5]
Numerically, this convention is aligned with the Robb-style “×10⁹” scale, but it remains non-product-specific and lacks full test-condition disclosure (temperature, humidity, instrument type).

## Comparison table, unit normalization, and flux estimates for wall-code 1 thickness

### Definitions and equations used

ASTM D1434 describes a manometric method that can yield gas transmission rate (GTR), permeance, and (for homogeneous materials) permeability, reported in SI units; it also notes tests are run with 0% moisture gas. [14]

For conversions and flux, the standard relationship used is: - Permeability coefficient: P
- Permeance: P/t
- Flux: J=P/t Δp (planar, steady-state) [15]

Unit conversions: - 1 Barrer=3.35×10−16 mol⋅m/m2⋅s⋅Pa [16]
- For STP volumetric reporting, this report uses the common “standard” convention 1 atm=101325 Pa at 0∘C. [17]

Note on “RTP vs STP Barrer”: Several silicone references (including Robb) report volumes referenced to 25 °C rather than 0 °C; converting “RTP cm³” to “STP cm³” reduces values by ~9%. Robb explicitly states his volumes are referenced to 1 atm and 25 °C. [18]
In the table below, “Barrer (STP adj)” applies that ~9% correction to enable consistent SI conversion.

### All found O₂/CO₂/N₂ values with normalized units and computed flux

Assumptions for flux: - wall thickness t = 0.79 mm (Saint‑Gobain wall-code 1) [7]
- Δp = 1 atm and 0.1 atm
- planar approximation (reasonable first order for thin walls; cylindrical correction discussed after table)

| Source | Gas | T (°C) | Original value | Original unit | Barrer_equiv_RTP (≈10×value) | Barrer (STP adj) | P_SI (mol·m)/(m²·s·Pa) | J @1 atm, 0.79mm (mol/m²/s) | J @1 atm (cm³ STP/m²/s) | J @0.1 atm, 0.79mm (mol/m²/s) | J @0.1 atm (cm³ STP/m²/s) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Robb 1968 Table 2 (filled dimethyl silicone rubber, 33% silica filler) | O2 | 25 | 60 | Pr×10^9 [cc gas (RTP)·cm]/(s·cm²·cmHg·ΔP) | 600 | 550 | 1.841e-13 | 2.362e-05 | 0.529 | 2.362e-06 | 0.053 |
| Robb 1968 Table 2 (filled dimethyl silicone rubber, 33% silica filler) | N2 | 25 | 28 | Pr×10^9 [cc gas (RTP)·cm]/(s·cm²·cmHg·ΔP) | 280 | 257 | 8.593e-14 | 1.102e-05 | 0.247 | 1.102e-06 | 0.025 |
| Robb 1968 Table 2 (filled dimethyl silicone rubber, 33% silica filler) | CO2 | 25 | 325 | Pr×10^9 [cc gas (RTP)·cm]/(s·cm²·cmHg·ΔP) | 3250 | 2977 | 9.975e-13 | 1.279e-04 | 2.868 | 1.279e-05 | 0.287 |
| Robb 1968 Table 5 (control, 0 roentgens) | O2 | 25 | 57 | Pr×10^9 [cc gas (RTP)·cm]/(s·cm²·cmHg·ΔP) | 570 | 522 | 1.749e-13 | 2.244e-05 | 0.503 | 2.244e-06 | 0.050 |
| Robb 1968 Table 5 (control, 0 roentgens) | N2 | 25 | 27 | Pr×10^9 [cc gas (RTP)·cm]/(s·cm²·cmHg·ΔP) | 270 | 247 | 8.287e-14 | 1.063e-05 | 0.238 | 1.063e-06 | 0.024 |
| Robb 1968 Table 5 (control, 0 roentgens) | CO2 | 25 | 285 | Pr×10^9 [cc gas (RTP)·cm]/(s·cm²·cmHg·ΔP) | 2850 | 2611 | 8.747e-13 | 1.122e-04 | 2.516 | 1.122e-05 | 0.252 |
| Limitless Shielding PDMS membrane tube (platinum-cure) – website values | O2 |  | 50 | "barrers" 10^-9 [cc gas (RTP)·cm]/(s·cm²·cmHg·ΔP) | 500 | 458 | 1.534e-13 | 1.969e-05 | 0.441 | 1.969e-06 | 0.044 |
| Limitless Shielding PDMS membrane tube (platinum-cure) – website values | N2 |  | 25 | "barrers" 10^-9 [cc gas (RTP)·cm]/(s·cm²·cmHg·ΔP) | 250 | 229 | 7.673e-14 | 9.841e-06 | 0.221 | 9.841e-07 | 0.022 |
| Limitless Shielding PDMS membrane tube (platinum-cure) – website values | CO2 |  | 270 | "barrers" 10^-9 [cc gas (RTP)·cm]/(s·cm²·cmHg·ΔP) | 2700 | 2474 | 8.286e-13 | 1.063e-04 | 2.382 | 1.063e-05 | 0.238 |
| Zhang & Cloud (Arlon) 2006 paper Table 2 (cites ref [9]) | O2 |  | 62 | Permeability×10^9 [cm³·cm]/(s·cm²·cmHg) | 620 | 568 | 1.903e-13 | 2.441e-05 | 0.547 | 2.441e-06 | 0.055 |
| Zhang & Cloud (Arlon) 2006 paper Table 2 (cites ref [9]) | N2 |  | 28 | Permeability×10^9 [cm³·cm]/(s·cm²·cmHg) | 280 | 257 | 8.593e-14 | 1.102e-05 | 0.247 | 1.102e-06 | 0.025 |
| Zhang & Cloud (Arlon) 2006 paper Table 2 (cites ref [9]) | CO2 |  | 323 | Permeability×10^9 [cm³·cm]/(s·cm²·cmHg) | 3230 | 2959 | 9.913e-13 | 1.271e-04 | 2.850 | 1.271e-05 | 0.285 |

Key takeaway: Across multiple sources, the consistent ranking is CO₂ ≫ O₂ > N₂, with CO₂ typically ~10× helium and ~5–12× O₂ in silicone rubbers, driven primarily by solubility plus high diffusion in the siloxane backbone. [19]

xychart-beta
  title "Representative permeability (Barrer-equivalent) from Robb 1968 Table 2"
  x-axis ["N2","O2","CO2"]
  y-axis "Permeability (Barrer-equivalent)" 0 --> 3500
  bar [280, 600, 3250]

### Geometry note: tubing vs flat-film approximation

For a first estimate, using J=P/tΔp with t=0.79 mm is reasonable. [20]
If you need a more exact cylindrical-wall calculation for total mol/s through a length of tube, use the radial steady-state form:

n=2πL P Δplnro/ri

and define whether Δp is across the inner/outer partial pressures of each gas species. (This becomes material-impacting when wall thickness is not “thin” relative to radius.)

## Standards context and recommended next steps

### Standards that map to what you need

ASTM D1434 provides a manometric method for gas transmission through films/sheeting and can yield permeance and permeability in SI units (and notes sensitivity to procedure/lab and moisture effects). [14]
For non-film materials and broader gas sets, laboratories commonly reference ISO/DIN families such as ISO 15105 (manometric gas transmission through films/sheetings) and rubber-focused methods (e.g., DIN routes), and the Mecadi lab listing explicitly offers O₂, CO₂ and N₂ permeability testing “according to ISO 15105.” [21]

### Practical next steps to obtain ULTRA‑C‑125‑1F/1H specific values

Request from Saint‑Gobain (most likely path): - Ask Saint‑Gobain Life Sciences for any measured O₂/CO₂/N₂ permeability for Sani‑Tech® Ultra‑C (and specifically the ULTRA‑C‑125 wall code 1 compound), and request test conditions + method (ASTM D1434 / DIN 53536 / ISO 15105 or equivalent), because the public site indicates dossiers/validation guides are gated. [22]

If data cannot be provided or must be generated: - Commission third‑party testing with an explicit protocol and reporting format: - Report permeability coefficient P, not just transmission rate, and specify temperature, humidity, gas purity, and pressure range. - For tubing, decide whether to test flat plaques cut from molded material (cleaner interpretation) or as-extruded tubing wall (more application-realistic). ASTM D1434 is film-focused; a lab may adapt fixtures or recommend a rubber-permeability method depending on sample geometry. [23]

Gap statement (what is still missing for ULTRA-C-125): - No public source located that ties Saint‑Gobain ULTRA‑C formulation to a measured O₂/CO₂/N₂ permeability coefficient with full test report metadata (method version, humidity control, specimen conditioning, uncertainty). [24]

[1] [6] [7] [8] [9] [24] https://www.biopharm.saint-gobain.com/sites/hps-mac3-lifesciences-bioprocess/files/2022-09/Sani-Tech%20Ultra-Silicone%20Tubing-datasheet.pdf

https://www.biopharm.saint-gobain.com/sites/hps-mac3-lifesciences-bioprocess/files/2022-09/Sani-Tech%20Ultra-Silicone%20Tubing-datasheet.pdf

[2] [22] Sani-Tech® Ultra Silicone Tubing | Bioprocess Solutions

https://www.biopharm.saint-gobain.com/components/fluid-transfer/sani-tech-ultra-silicone-tubing

[3] [4] [10] [11] [18] [19] https://gwern.net/doc/technology/1968-robb.pdf

https://gwern.net/doc/technology/1968-robb.pdf

[5] Gas Permeable PDMS Membrane Tube - Limitless Shielding

https://limitless-shielding.com/pdms-membrane-tube/

[12] [13] Microsoft Word - D116, Haibing Zhang, et al.doc

https://cdn.thomasnet.com/kc/2650/doc/0000102048_70_83571.pdf

[14] [23] https://store.astm.org/standards/d1434

https://store.astm.org/standards/d1434

[15] [16] [20] The Impact of Various Natural Gas Contaminant Exposures on CO

https://www.mdpi.com/2077-0375/10/11/324?utm_source=chatgpt.com

[17] Pressure and Gas Flow Unit Conversions | NIST

https://www.nist.gov/pml/owm/metric-si/unit-conversion/pressure-and-gas-flow-unit-conversions?utm_source=chatgpt.com

[21] https://www.mecadi.com/en/permeation-measurement/

https://www.mecadi.com/en/permeation-measurement/